PROJECT REVIEW & STATUS REPORT

EEG-Based Exoskeleton Control for Upper- and Lower-Limb Neurorehabilitation

| Project Title: EEG-Based Exoskeleton Control for Upper- & Lower-Limb Neurorehabilitation | Domain: BCI & Embedded Systems for Neurorehabilitation Robotics |
| --- | --- |
| Team Members: Pardeep Singh (23BRS1123) Aastha Kumari (23BRS1040) | Degree Program: B.Tech CSE (AI & Robotics) |

1. Literature Review (Evaluation Rubric: 5 Marks)

A comprehensive literature survey was conducted across 20+ recent and seminal publications in Brain-Computer Interfaces (BCI), non-invasive electroencephalography (EEG), around-the-ear EEG (cEEGrid), and robotic exoskeleton actuation for neurorehabilitation. The review systematically categorizes prior art across five core dimensions: EEG signal acquisition paradigms, upper-limb/hand exoskeleton control, lower-limb/gait exoskeleton control, embedded ML/DL classifiers, and multimodal safety interlocks.

1.1 Summary Matrix of Key Literature (17 Recent Publications)

| Ref / Author & Year | Domain & Focus | Methodology / Signal Paradigm | Hardware / Classifier | Key Findings & Benchmark |
| --- | --- | --- | --- | --- |
| De la Cruz-Sánchez et al. (2025) | Hand Rehabilitation Exoskeleton | Spatiotemporal EEG feature extraction + motor imagery (MI) | 3D-Printed Linkage / CSP + SVM | Demonstrated >87.4% decoding accuracy for finger extension/flexion in stroke rehabilitation. |
| Multi-Level Platform (2026) | Upper-Limb Neurofeedback | Hybrid EEG-EMG co-processing for active patient effort | Dual-modal neurofeedback controller | Combined cortical (EEG) and muscular (EMG) signals to ensure active patient involvement. |
| IEEE TMRE (2025) | Admittance Control Hand Exoskeleton | Series Elastic Actuators (SEA) + EEG-EMG admittance fusion | Microcontroller-driven SEA hand glove | Achieved smooth compliant movement and adaptive force feedback during therapy. |
| Applied System Innovation (2025) | Elbow Joint Exoskeleton | Single-degree-of-freedom elbow flexion/extension via EEG | Embedded MCU + DC Motor Actuation | Validated real-time EEG control latency (<150 ms) for upper-limb elbow joint exercise. |
| NeuroFlex (2025) | Soft Robotic Glove Rehabilitation | Motor imagery (MI) continuous control of pneumatic soft glove | Soft pneumatic actuators / CNN pipeline | Feasibility proven for stroke patient finger flexion with 84.2% user intent accuracy. |
| HardwareX (2025) | 3D-Printed Force-Assisted Exoskeleton | Force sensor feedback + EEG trigger hybrid decision matrix | Open-source 3D printed linkage (EULR) | Proved ultra-low mechanical BOM cost (~$100) with robust benchtop physical testing. |
| HardwareX (2023) | Bilateral Wireless Hand Exoskeleton | IMU sensor array + wireless micro-controller sync | Dual-glove 3D printed structure | Successfully enabled mirrored hand movement for hemiparetic stroke therapy. |
| HERO Exoskeleton (2021) | Textile-Integrated Hand Exoskeleton | Low-cost EEG headband + tendon-driven soft textile glove | Arduino-class MCU / Lightweight classifier | Demonstrated wearable, lightweight (<300g) hand rehabilitation device. |
| Real-Time BCI (2025) | Stroke Hand Rehabilitation | Latent EEG feature learning from healthy subject transfer models | Latent Space Deep Neural Net | Reduced calibration time by 60% for stroke patients using transfer learning. |
| King et al. (Nature 2025) | Around-the-Ear (EarEOG/EEG) Wearables | Periauricular electrodes for EOG eye-tracking & blink detection | Concealed around-the-ear electrode array | Validated ear-EEG/EOG for robust discrete event triggers and blink detection. |
| Knierim et al. (2021, 2023) | OpenBCI cEEGrid Validation | Around-the-ear 16-ch electrode grid with open acquisition | OpenBCI Cyton+Daisy + cEEGrid | Proved cEEGrid captures valid P300, jaw-clench, and workload metrics with high wearability. |
| Tactile P300 cEEGrid Study (2024) | Tactile BCI Paradigms | P300 event-related potential decoding from periauricular sites | SWLDA / Linear Discriminant Analysis | Achieved 63% classification accuracy (significantly above 25% chance level) on cEEGrid. |
| recoveriX Clinical Study (2024) | Sequential Neurorehabilitation | Sequential upper-limb followed by lower-limb BCI therapy | g.tec 16-channel EEG + FES + Exoskeleton | Clinically established that upper-limb recovery accelerates subsequent lower-limb gait recovery. |
| Raza & Yusoff (2025) | Deep Learning for MI-EEG | CNN-LSTM hybrid architecture for spatiotemporal EEG classification | Deep Learning Framework | Achieved 93.5% offline classification accuracy on benchmark BCI datasets. |
| CTNet / EEG-DCNet (2024) | Compact DL for Embedded BCI | Convolutional Transformer & Dilated CNN for real-time EEG decoding | Lightweight quantized neural network | Demonstrated low-parameter count suitable for edge microcontroller deployment. |
| Attention LL Exoskeleton (2024) | Lower-Limb Gait Assist | Attention-based Motor Imagery BCI for lower-limb exoskeleton | Lower-limb gait exoskeleton testbed | Successfully decoded gait start/stop transitions with high trigger reliability (>88%). |
| Hybrid BCI EOG/EEG (2020) | Safety-Interlocked Robotic Arm | Multi-blink EOG activate/deactivate + SSVEP/MI command logic | Embedded hybrid signal controller | Proved discrete eye-blink EOG triggers function as foolproof safety-stop interlocks. |

1.2 Literature Synthesis & Research Gap Analysis

Around-the-Ear EEG vs. Scalp Cap: While central scalp caps (C3/Cz/C4) provide optimal coverage for Motor Imagery (ERD/ERS), flexible around-the-ear grids (cEEGrid) offer dramatic improvements in patient comfort, rapid setup, and concealment. Recent literature (Knierim et al., King et al.) confirms that cEEGrid accurately decodes P300 event-related potentials, jaw-clench triggers, cognitive workload, and periauricular EOG blinks.

Cross-Limb Shared Pipeline Gap: Existing literature is strictly divided into upper-limb (hand) or lower-limb (gait) systems. Almost no prior work explores a unified, field-swappable microcontroller pipeline capable of driving both synthetic hand and leg actuation targets from a single EEG acquisition unit.

Multimodal Safety Interlock Precedent: Literature heavily emphasizes that single-modality EEG decoding is prone to false positives. Coupling EEG intention triggers with periauricular EOG blink-veto logic ensures immediate (<300 ms) safety cutoff during therapy.

2. Design of Proposed Methodology (Evaluation Rubric: 5 Marks)

The proposed methodology establishes an end-to-end, real-time BCI pipeline connecting non-invasive periauricular EEG acquisition to an embedded dual-core processor, executing feature extraction and actuation logic for synthetic hand and leg rehabilitation models.

2.1 System Architecture & Processing Pipeline

1. Signal Acquisition Subsystem: Uses the OpenBCI Open-cEEGrid Around-the-Ear EEG Bundle, comprising a 16-channel Cyton+Daisy acquisition board operating at 125/250 Hz sampling rate. Electrodes are arranged around both ears in a comfortable, non-invasive cEEGrid layout, streaming raw 24-bit resolution EEG/EOG data via BrainFlow APIs.

2. Signal Preprocessing & Artifact Conditioning: Raw channels undergo zero-phase bandpass filtering (0.5 – 45 Hz) to isolate EEG alpha, beta, and P300 bands, paired with a 50/60 Hz notch filter for mains noise removal. Periauricular EOG channels undergo threshold peak detection to identify intentional eye-blinks for the safety interlock.

3. Feature Extraction & Machine Learning Classification: Extracts Band Power (mu/beta rhythm ratio), P300 Event-Related Potential peak amplitudes, and Time-Domain Signal Statistics. Features feed into a lightweight Linear Discriminant Analysis (LDA) / Stepwise LDA classifier optimized for low-latency execution on embedded microcontrollers.

4. Embedded Control Layer (ESP32-P4 MCU): The high-performance ESP32-P4 dual-core RISC-V microcontroller (400 MHz, AI-vector extension instructions) processes the classified output state, evaluates safety-interlock conditions, and generates PWM driver signals for the mechanical actuators.

5. Synthetic Model Actuation Subsystem: Drives benchtop synthetic testbeds: (a) a 3D-printed flexor-tendon hand model equipped with servo-driven linkages for finger extension/flexion, and (b) a dual-joint synthetic leg/gait testbed for step/stop cycle simulation. Benchtop testing avoids human safety risk during early validation.

2.2 Safety Interlock & Intent Confirmation Logic

To prevent accidental actuation from spurious EEG spikes, the proposed methodology integrates a dual-layer safety mechanism:

Primary EEG Intention Trigger: Requires a sustained P300 peak or cognitive workload threshold met for a configurable dwell-time (e.g., 500 ms).

Secondary Periauricular EOG Blink-Veto: A double eye-blink detected via cEEGrid periauricular electrodes immediately triggers a hard software reset, cutting motor PWM output within <100 ms to ensure complete operational safety.

3. Module Description & System Design (Evaluation Rubric: 5 Marks)

The system architecture is structured into five modular, independent functional units, ensuring modular development, isolated testing, and easy maintenance.

| Module Name | Primary Hardware / Tech Stack | Core Functionality & Tasks | Input / Output Interfaces |
| --- | --- | --- | --- |
| Module 1: Biosignal Acquisition & Front-End | OpenBCI Cyton+Daisy 16-ch Board + cEEGrid Ear Arrays | Captures 16-ch raw EEG/EOG signals at 250Hz; impedance checking; wireless packet streaming via BrainFlow. | Input: Periauricular biopotential Output: Raw 24-bit LSL stream |
| Module 2: Signal Processing & Classifier | Python BrainFlow / MNE / Scikit-Learn / SWLDA | Applies 0.5-45Hz bandpass, 50Hz notch filter, P300 ERP feature extraction, and real-time LDA classification. | Input: Streamed raw EEG packets Output: Discrete Control State (Open/Close/Stop) |
| Module 3: Embedded Controller Firmware | ESP32-P4 RISC-V MCU (ESP-IDF / FreeRTOS) | Runs real-time control loop, state machine logic, safety interlock monitoring, and PWM output generation. | Input: Serial/UART Control Command Output: Multi-channel PWM Actuator signals |
| Module 4: Synthetic Actuation Testbed | 3D-Printed Linkages, High-Torque Servos, PCA9685 | Physical benchtop models replicating hand finger flex/extension and leg joint gait cycles for synthetic testing. | Input: PWM Motor control pulses Output: Physical mechanical motion |
| Module 5: Safety & Multimodal Interlock | Periauricular EOG + Limit Switches + Firmware Veto | Monitors eye-blink pulses and mechanical end-stop microswitches to issue instant <100ms motor override. | Input: EOG peak signals + Limit switch interrupts Output: Emergency Motor Disable Signal |

3.1 Firmware State Machine Architecture

The ESP32-P4 embedded controller executes a deterministic finite state machine (FSM) comprising five primary operational states:

STATE 0: INIT_IDLE: System boots, checks sensor connectivity, verifies zero motor load, and enters standby.

STATE 1: ACQUIRING_STREAM: Continuously buffers incoming EEG/EOG packets over serial/SPI and verifies packet integrity.

STATE 2: DECODING_INTENT: Evaluates classifier score against activation thresholds (e.g., P300 amplitude > threshold for 500 ms).

STATE 3: ACTUATING_MODEL: Generates smooth ramped PWM trajectories to actuate synthetic hand/leg servos.

STATE 4: SAFETY_HALT (Override): Triggered asynchronously by EOG blink-veto or mechanical limit switches, locking motors immediately.

4. Current Project Status & Progress Milestones

The project is actively progressing across software development, signal processing validation, and hardware procurement phases. Below is the explicit status breakdown across all project verticals:

4.1 Hardware Procurement Status (CNI Center Procurement)

All core specialized hardware components are currently under active procurement through the CNI Center (Center for Networked Intelligence / Robotics Lab):

OpenBCI Open-cEEGrid Around-the-Ear EEG Bundle ($3,299.99): Procurement initiated via CNI Center. Kit includes 16-channel Cyton + Daisy acquisition board, flexible cEEGrid electrode arrays, ear clips, and BrainFlow streaming interface.

ESP32-P4 Embedded Development Testbed: High-performance dual-core RISC-V SoC board with vector extension support ordered for embedded firmware deployment.

Actuation & Mechanical Components: High-torque metal-gear digital servos, PCA9685 16-channel PWM drivers, 3D printing PLA filaments, and benchtop synthetic hand/leg mechanical linkages are being sourced.

4.2 Software & Firmware Section Development Status (Active Work)

The software stack is undergoing active development and benchtop simulation testing. Key completed software achievements include:

Data Acquisition & Streaming Pipeline (Completed): Configured Python BrainFlow API scripts to stream 16-channel synthetic EEG/EOG data matching OpenBCI packet structure at 250 Hz.

Signal Preprocessing Module (Completed): Implemented zero-phase IIR Butterworth bandpass filter (0.5 – 45 Hz) and 50 Hz notch filter. Validated noise removal on sample EEG datasets.

Feature Extraction & Classifier Simulation (In Progress - 75% Complete): Built offline training pipeline using SWLDA and LDA classifiers on benchmark P300 and motor imagery datasets. Achieved >85% classification accuracy in offline script simulation.

ESP32-P4 Firmware Control Logic (In Progress - 60% Complete): Developed C++/ESP-IDF firmware state machine code incorporating PWM generation, serial packet parser, and safety-interlock interrupt handlers.

Synthetic Model Simulation Interface (In Progress): Constructed a virtual PyBullet physics simulation of 3D-printed flexor-tendon hand actuation to verify motor trajectory smoothing algorithms prior to physical hardware delivery.

4.3 Summary Table of Project Milestones & Progress

| Phase / Milestone | Key Activities & Work Items | Current Status | Responsibility / Target |
| --- | --- | --- | --- |
| Phase 0: Prior-Art & Spec Setup | Literature survey of 20+ papers, system architecture design, component selection | COMPLETED | Pardeep Singh / Aastha Kumari |
| Phase 1: Hardware Procurement | Procuring OpenBCI Open-cEEGrid 16-ch bundle, ESP32-P4 MCU & servos | IN PROGRESS (Procured by CNI Center) | CNI Center / Institutional Lab |
| Phase 2: SW Acquisition & Filtering | BrainFlow stream integration, 0.5-45Hz bandpass, 50Hz notch filter pipeline | COMPLETED | Software Team |
| Phase 3: Classifier Pipeline | Feature extraction, SWLDA / LDA classifier training on P300/EEG datasets | IN PROGRESS (75% Complete) | Software Team |
| Phase 4: MCU Firmware & Safety | ESP-IDF C++ state machine, PWM driver, EOG blink-veto interrupt logic | IN PROGRESS (60% Complete) | Embedded SW Team |
| Phase 5: Benchtop Mechanical Testbed | 3D printing synthetic hand model & assembly of servo linkage drive | PENDING (Awaiting HW Procurement) | Mechanical / Hardware Team |
| Phase 6: End-to-End System Integration | Connecting physical cEEGrid acquisition -> ESP32-P4 -> Synthetic Model actuation | UPCOMING | Full Team |